Decision-threshold and detection-limit calculator for high-resolution gamma-ray spectrometry

Description and explanation of the method and user’s guide

Jožef Stefan Institute, Jamova cesta 39, Ljubljana, Slovenia

1.  Summary

The characteristic limits, i.e., the decision threshold [y*, Eq. (6)] and the detection limit [y#, Eq. (7)], are the key quantities that describe the sensitivity of a measurement. They are given by the uncertainties that contribute to the measurement result [y, Eq. (1)] and by the measurement result itself. Besides the uncertainties arising from the calculation of the net number of counts in a peak [nn, Eq. (A2)], which presents the net indication and are specific for gamma-ray spectrometric measurements, the uncertainty of the conversion factor (W), originating in the process of sampling and sample preparation, also contributes to the characteristic limits. To make the calculations more surveyable, the characteristic limits are calculated in two steps: in the first step the uncertainty of the net number of counts corresponding to the characteristic limits is calculated, and in the second, the characteristic limits themselves.

The easiest way to calculate the characteristic limits is to start from the peak-analysis data, i.e., the number of counts in the peak (np) and its uncertainty, reported by the peak-analysis programs. Additional data needed for the calculation are the gross number of counts in the peak region or the region of interest (ng), which must be retrieved from the spectrum, and the data on the possible peaked background (nB), which can be retrieved from the background-subtraction and interference-correction reports.

To arrive at the uncertainty of the net number of counts, the uncertainties of the peak area and the uncertainty of the number of counts in the peaked background must be considered. While the peak area and the number of counts in the peaked background are not correlated, because they are obtained from different measurements, the peak area and the number of counts in the continuous background under the peak [n0, (Eq. A1)] might or might not be correlated, depending on the method of evaluation. If the peak area is calculated with the Region of Interest (ROI) method, the peak area and the number of counts in the continuum are not correlated, because they are obtained from non-overlapping spectral regions. If the peak area is calculated with the Least Squares (LSQ) method, the correlation exists, since the peak area and the number of counts in the continuum are obtained from the same spectral region. To cover most cases that can be encountered in the peak analysis of the gamma-ray spectra, several methods for the calculation of the characteristic limits are implemented in the Decision threshold and detection limit Calculator (the Calculator) that are applicable to the following situations:

- isolated peaks analyzed by the ROI or LSQ methods,

- overlapping peaks evaluated by the LSQ method,

- multi-gamma-ray emitters.

If the indication was not recognized in the peak-location step of the peak analysis, no peak-analysis data are available for the calculation and the characteristic limits must be calculated by the ROI method, assuming zero for the peak area and an uncertainty calculated from the gross number of counts in the peak region.

2. Introduction

The standard ISO 11929-1:2019(E) [1] requires reporting the decision threshold and detection limit for all the measurands of interest, regardless of whether their indications were recognized in the spectrum or not. It is relatively easy to arrive at the decision threshold and detection limit in the absence of the indication, since in this case the detection limit and decision threshold are given only by the uncertainty of the background, where the indication is expected to appear. If the indication is present in the spectrum, then the decision threshold and the detection limit can depend on it, since a larger indication hides the background more effectively. Therefore, the information about it is less accurate and consequently its uncertainty is larger. To arrive at the decision threshold and the detection limit in this case it is necessary to assess the uncertainty of the indication in a measurement that was not performed.

The standard ISO 11929 [1] requires establishing the uncertainty of the observed value y as a function of its value and then, in order to calculate the decision threshold, evaluating u(y=0) and solving the equation y=k1-αu(y=0)+k1-βu(y) to calculate the detection limit. The Calculator computes the decision threshold and the detection limit from the peak parameters, which are commonly listed in the peak-analysis reports and from the total number of counts registered in the region of the spectrum where the indication resides.

Usually, in gamma-ray spectrometry two methods of peak-area calculation are used: The Region of Interest (ROI) method and the Least Squares (LSQ) method. Since the suppositions of the methods are different, the methods lead to different expressions for u(y).

The Calculator has different codes for evaluating the decision threshold and the detection limit for:

- isolated peaks evaluated with the ROI method, Appendix A,

- isolated peaks evaluated with the LSQ method [2, 3],

- two overlapping peaks, evaluated by the LSQ method, one of the peaks is treated as the indication, the other as belonging to the background [2, 3],

- two overlapping peaks, evaluated by the LSQ method, where one peak is treated as the indication with properties depending on both peaks and the energy distance between them [4].

For convenience, to enable the calculation of the decision thresholds and detection limits for measurands not identified in the spectrum, the Calculator also has a procedure for calculating the decision thresholds and detection limits with the ROI method, starting from the channel contents. It also has a procedure for calculating the decision thresholds and detection limits for multi-gamma-ray emitters.

3. Calculation of the decision thresholds and detection limits

3.1 General

It is a common assumption in the derivations of all the decision-threshold and detection-limit calculations that the height of the continuous background, on which the indication is superimposed, is independent of the intensity of the indication. However, its uncertainty can depend on its intensity since a larger indication hides the background more effectively. It should be noted that the assumption of the independence of the continuous background, when the contribution of the small-angle scattering of gamma-rays to the continuous background at the low-energy side of the indication is large or when the influence of the pile-up effect at the high-energy side of the indication is large, might not be justified.

In contrast to the uncertainty of the continuous background, the uncertainty of the peaked background cannot depend on the intensity of the indication. The peaked background originates in the spectrometer background, the activity of the blank sample and the contributions of the spectral interferences. The contribution of the first two sources is obtained from independent measurements, whereas the contribution of the interferences is calculated from statistically independent data, obtained in other spectral regions of the same spectrum. Therefore, the peaked background is not statistically correlated with the intensity of the indication.

The total number of counts in the peak region carries information about the continuous background. The width of the peak region, used for the calculations of the decision threshold and the detection limit, depends on the expressiveness of the peak and amounts to 1.25 FWHM for unexpressive peaks and 2.5 FWHM for expressive peaks [5]. If the tails of the peak are not excessively long, then for peaks dominating the spectrum it should not exceed 3 FWHM.

Since the uncertainty of the gross number of counts (given by the Poisson statistic) and the uncertainty of the indication at the observed value of the measurand are known, it is possible to conclude from the measurement model and the law of propagation of uncertainties, by assessing or supposing the values of the correlation coefficients, about the uncertainty of the continuous background and consequently also about the uncertainty of the indication as a function of its value.

To calculate the decision thresholds and the detection limits for gamma-ray spectrometric measurements, two approaches are possible:

1. Directly, from channel contents. This approach is presented in [6]. The spectrum is unfolded using the LSQ method. The uncertainty of the measurand as a function of its value is calculated numerically from the model and the detection limit is calculated by iteration. Alternatively, in the case of isolated peaks, it can be analyzed using the ROI method. The decision threshold and the detection limit are calculated from the uncertainty of the continuous background in the spectral region where the peak resides [7, 8]. If the presence of the signal due to the measurand was not recognized, then it can be assumed that no peak is present at the energy where the indication is expected to occur. Then, the decision threshold and the detection limit can be calculated from the uncertainty of the sum of channel contents within the region with a width corresponding to the width of the unexpressive peaks. This assumption is used in most software packages intended for the analysis of high-resolution, gamma-ray spectra, regardless of the sensitivity of the peak location in the peak analysis.

2. From the peak-analysis data and the total number of counts in the energy region where the indication appears. For calculating the decision threshold and detection limit it is not necessary to unfold the spectrum since in the case of the recognition of the signal due to the measurand, the unfolding of the spectrum has already been accomplished by the peak analysis. The uncertainty of the measurand as a function of its value is obtained from the model by the law of propagation of uncertainties in analytical form. Since all the data necessary to construct this function might not be available in the peak-analysis reports, values for some quantities (correlation coefficients) are assumed or assessed.

In gamma-ray spectrometry, peak areas represent the indications; therefore, the use of peak-analysis results in the calculations of the decision threshold and detection limit is preferable to the use of the raw spectral data, i.e., the channel contents. Then, the quality of these calculations reflects the quality of the peak analysis used for determining the measurement results: overestimated peak-area uncertainties increase the uncertainty of the continuous background, while underestimated uncertainties decrease it and can even prevent the calculation of the background uncertainty. By using peak-analysis results, the decision threshold and detection limit are compatible with them.

For indications recognized in the spectrum, the Calculator uses the peak-analysis results as input quantities.

If the presence of the measurand (the indication) was not recognized in the spectrum, no peak-analysis results are available. In this case, three possibilities for the calculation of the decision threshold and the detection limit exist:

-  calculation from the channel contents using the ROI method (implemented in the Calculator located on the sheet designated by the ROImethod), or alternatively,

-  the calculation can be performed with the peak-area uncertainty calculated from the total number of counts in the peak region, by calculating the uncertainty of the peaked background (if present) and by supposing a value of zero for the net and background peak areas (implemented in the Calculator located on the sheet designated by ROIpeak),

-  inserting interactively a peak at the energy where the signal is expected with an appropriate software tool (if available) and calculating the decision threshold and the detection limit from the peak data.

3.2 Applications of the ROI and LSQ methods to specific cases

Usually, in gamma-ray spectrometry two methods for the peak-area calculation are used: The Region Of Interest (ROI) method and the Least Squares (LSQ) method. The basic difference between the two methods is that with the LSQ method the spectrum within the energy region of interest is unfolded into its components: the peaks and the continuous background. With the ROI method, the spectrum is decomposed into one isolated peak and the continuous background using non-overlapping spectral regions, one of them comprising the peak and the others the continuous background only. From these regions the continuous background under the peak is calculated. The ROI method might not be applicable to the peaks that are strongly dominating the spectrum. Such peaks can exhibit tails, which prevent a reliable calculation of the continuous background in the energy regions adjacent to the peak. It is not possible to evaluate overlapping peaks with the ROI method; here, the LSQ method must be used.

The different methods imply different correlation coefficients between the numbers of counts registered in the peaks and in the continuous background. It is easy to see that in the case of the ROI method, the correlation coefficient between the total number of counts in the peak region and the number of counts in the regions comprising the continuous background only is zero. It follows that the uncertainty of the height of the continuous background under the peak is given by the number of counts in the regions adjacent to the peak region. Therefore, the decision threshold and the detection limit depend on the intensity of the indication only through the width of the peak region.

On the other hand, when decomposing the spectrum into a peak and into the continuous background with the LSQ method, the correlation coefficient between the number of counts in a peak and the number of counts in the continuous background is less than zero. Namely, the total number of counts in the whole energy region being decomposed is fixed, since it is equal to the sum of the channel contents. Therefore, when increasing the peak area for a specified number of counts, the number of counts in the continuous background must decrease by the same number. The value of the correlation coefficient depends on the width of the energy region being decomposed. For regions much wider than the peak width, the background is determined with the spectrum outside the peak region, and because of that the correlation coefficient assumes here a small negative value [9]. However, in the limit of narrow energy regions comprising the peak region only, and this is the case encountered in the peak analyses, the value of the correlation coefficient approaches –1. This value is assumed in the calculations of the decision thresholds and detection limits for isolated peaks and overlapping peaks, where the overlapping peak is treated as a part of the background.

In the case of decomposing the spectrum into two overlapping peaks and in the continuous background the correlation coefficients can assume different values, depending on the spacing between the peaks and on the width of the spectral region to be decomposed [4,10]. Here, in the calculation of the decision thresholds and detection limits with the LSQ method, it is supposed that the continuous background is constant within the whole spectral region that is unfolded.

4. The method of calculation

The total number of counts in the peak region carries the information about the continuous background. Since the uncertainty about the gross number of counts is known from the Poisson statistic and the uncertainty of the indication at the observed value of the measurand is known from the peak analysis, it is possible to conclude from the law of the propagation of uncertainties [11], by assessing or supposing the values of the correlation coefficients, about the uncertainty of the continuous background and consequently about the uncertainty of the indication as a function of its value.

The decision thresholds and detection limits are calculated in two steps. In the first step, the numbers of counts corresponding to the detection limit and the decision threshold are calculated, and in the second step, the decision thresholds and detection limits themselves.

The relation between the net number of counts (the net indication) nn and the observed value of the measurand y is

y = W nn  ,									(1)

where W denotes the conversion factor, comprising all the data describing the measurand, the sample and detector properties relevant to the determination of the observed value of the measurand.

The corresponding relation among the uncertainties is

u2(y) = y2 [urel2(W) + urel2(nn)]  ,						(2)

where urel denotes the relative uncertainty.

The decision threshold is defined as

y* = k1-α u(y=0)								(3)

and the detection limit as

y# = y* + k1-β u(y=y#)   ,							(4)

where the uncertainty from Eq. (2) at the detection limit is

u2(y=y#) = y#2 [urel2(W) + u2(nn=nn#)/nn2#]   .					(5)

Since

y* = W nn*									(6)

and

y# = W nn#   ,									(7)

it follows from Eq. (4)

(nn# - nn*)2 = k1-β2 [urel2(W) nn#2 + u2(nn=nn#)]   .				(8)

By deriving u(nn) from the law of propagation of uncertainties from the total number of counts and the peak-analysis results, from Eq. (8) it is possible to extract the unknown nn#. To solve the equation analytically, u2(nn) must be approximated with a polynomial of the second order and inserted in Eq. (8). Then, nn# can be calculated as the solution of the quadratic equation

A n#2 +B n# + C = 0,								(9)

where the coefficients A, B and C are given by the decision threshold, by the relative uncertainty of the conversion factor, by the results of the peak analysis, by their uncertainties and by the correlation coefficients among the results of the peak analysis.

To illustrate the procedure for the derivation of the decision threshold and the detection limit, in appendix A the derivation for the case of the ROI method, which resembles counting measurements, is presented. The case of an isolated peak, analyzed with the LSQ method, is described in [2] and [3], the case of overlapping peaks, analyzed by the LSQ method, where one peak is treated as part of the background, is described in [2] and [3] as well as the case of two overlapping peaks in [4].

The decision thresholds and the detection limits for multi-gamma-ray emitters are determined from the decision thresholds and detection limits calculated at different energies. It is not possible to perform the calculations from the corresponding indications occurring at different energies, because the relative uncertainties of the conversion factors can vary with the energy. The method of calculation is described in [2] and [3].

5. Description of the Calculator

The Calculator is an Excel file. To achieve a better survey over the calculations, the calculations are separated into nine sheets. Each sheet performs the calculation of a specific task. Five of the sheets perform calculations of the decision thresholds and detection limits (ROImethod, ROIpeak, LSQisolated, LSQoverlapping, LSQnp1np2), three of them are used for assessing the value of the correlation coefficient describing the correlation of the number of counts registered in two overlapping peaks (rCorrection, correlcoef, correlcoefB) and one sheet calculates the indication corresponding to the decision threshold and the detection limit for multi-gamma-ray emitters (MultiGamma). To calculate the indications corresponding to the decision threshold and the detection limit with the sheet LSQnp1np2, the input data in the three sheets performing the calculation of the correlation coefficients must be updated.

The analysis of the spectral shape of the overlapping peaks is performed by unfolding the spectrum into the two peaks and the continuous background using the LSQ method. Since the peak-analyzing programs do not report the correlation coefficients, it is necessary to reconstruct the spectrum shape from the total number of counts in the region of interest, the positions of the peaks, their FWHM values and the peak areas. The correlation coefficient between the peak areas is then assessed from the reconstructed spectrum. The Calculator computes the correlation coefficient for two extreme cases: for the case of large peaks on a small background and for the case of small peaks on a large background [4]. To differentiate between the large and small peaks the value of the peak-significance parameter is used, which in the case of small peaks assumes the value of s0 = np/n01/2 [12]. When a specific doublet is analyzed, the value of the correlation coefficient for expressive peaks is used in the calculation of the decision threshold and the detection limit if the peak significance parameter s0, calculated with the mean of both peak areas, exceeds the boundary value, which is set to s0 = 4√3. This value of the peak-significance parameter corresponds to the relative uncertainty of the sum of both peak areas uR(np1+np2) between 20% and 40%, depending on the peak areas. These relative uncertainties broadly resemble the relative uncertainty of the number of counts, corresponding to the detection limit [12]. Otherwise, the correlation coefficient for unexpressive peaks is used.

The cells in the Excel sheets appear in three colors: in green are the cells containing the input data, the red cells contain intermediate results and descriptions of the contents of the cells. The end results calculated in each sheet and the decision thresholds and detection limits are in yellow. Some comments, explaining the enquiring user values of correlation coefficients calculated under different suppositions, are not colored.

List of symbols used in the Excel file:

kp	quantile of the standardized normal distribution for the probability p

alpha	probability of making an error of the first kind

beta	probability of making an error of the second kind

FWHM	full width of the peak at half of its maximum

DeltaCh1 energy width of one channel

h	peak width in channels

Ei	energy

Ep	energy, where a peak is expected to be registered

Delta E energy difference

sigma	parameter of the normal distribution

Nb	number of channels in the background regions

Np	number of channels in the peak region

n0, nb	number of counts in the continuous background

nB	number of counts in the peaked background

np 	number of counts in the peak

nn	net number of counts in the peak

ng	gross number of counts in a spectral region comprising the indication

nn*	net number of counts corresponding to the decision threshold

nn#	net number of counts corresponding to the detection limit

W	conversion factor

A,B,C	coefficients of the quadratic equation

y(E)	primary measurement result obtained at energy E

y*	decision threshold

yC*	decision threshold of a multi-gamma-ray emitter

y#	detection limit

yC#	detection limit of a multi-gamma-ray emitter

r(np1,np2) correlation coefficient between the quantities np1 and np2

xi	matrix element of the design matrix P

Ni	matrix element of the variance-covariance matrix S

Nij	matrix element of the normal matrix N

M	parameter describing the expressiveness of a peak

Remarks:

u(a)	uncertainty of the quantity a

u(a=x)	uncertainty of the quantity a if it assumes a specific value x

urel(a)	relative uncertainty of the quantity a

6. Description of the example files

With the Calculator file comes an example spectrum file ExampleCNF, containing a gamma-ray spectrum in Canberra’s .CNF format, and a file comprising the description of its peak analysis, the peak-analysis results and the characteristic limits following from the peak-analysis results (Tables 1, 2 and 3) with a detailed description of how to manage the input quantities to the Calculator and its results. To calculate the characteristic limits for isolated peaks the peak-analysis results should be entered (Tables 1 and 2). If the indication overlaps with a neighboring peak, the peak-analysis data for the indication are entered into the cells np1 and u(np1) of the Excel sheet and the peak-analysis data of the overlapping peak into the cells np2 and u(np2). The energy-specific decision thresholds and detection limits for the multi-gamma-ray emitter Cs-134 are presented the in Table 3 for a calculation of the common characteristic limits for Cs-134.

References

1. ISO, ISO 11929-1:2019(E). Determination of the characteristic limits (decision threshold, detection limit and limits of the coverage interval) for measurements of ionizing radiation – Fundamentals and application, ISO,

Geneva, Switzerland.

2. Korun, M, Vodenik, B, Zorko, B, 2016. Calculation of the decision thresholds for radionuclides identified in gamma-ray spectra by post-processing peak analysis results, Nucl. Instr. and Meth. A 813, 102.

3. Korun, M, Vodenik, B, Zorko, B, 2018. Calculation of the detection limits for radionuclides identified in gamma-ray spectra based on post-processing peak analysis results, Appl. Radiat. Isot. 133, 22.

4. Korun, M.,  Petrovič, T, Vodenik, B, Zorko, B,  Calculation of the decision thresholds and detection limits in the case of overlapping peaks accruing in gamma-ray spectra by post-processing peak analysis results, submitted for publication.

5 ISO, ISO 11929-2010(E). Determination of the characteristic limits (decision threshold, detection limit and limits of the coverage interval) for measurements of ionizing radiation – Fundamentals and application, ISO, Geneva, Switzerland.

6. ISO, ISO 11929-3:2019(E). Determination of the characteristic limits (decision threshold, detection limit and limits of the coverage interval) for measurements of ionizing radiation – Applications to unfolding methods, ISO, Geneva, Switzerland

7. Currie, L A., Limits of qualitative detection and quantitative determination. Application to radiochemistry. Annal. Chem, 40, 586 – 593, 1968.

8.Gilmore, G., Gamma-ray spectrometry, 2011, John Willey & Sons, Chichester, England

9. Korun, M, Petrovič T, Vodenik, B, Zorko, B, 2020. Empirical determination of the correlation coefficient between the number of counts in a peak in a gamma-ray spectrum and the number of counts in the continuum where the peak is superimposed, Appl. Radiat. Isot. 158, 109063.

10. Korun, M, Vodenik, B, Zorko, B. 2016. Calculation of the correlation coefficients between the numbers of counts (peak areas and backgrounds) obtained by gamma-ray spectra, Appl. Radiat. Isot. 118, 1.

11. ISO, 2009. Guide to the Expression of uncertainty in measurement, ISO Switzerland.

12. IAEA, 2017. Determination and interpretation of characteristic limits for radioactivity measurements, IAEA/AQ/48, Vienna, Austria.

Appendix A. Derivation of the decision threshold and detection limit for the ROI method starting from the peak-analysis results for the case of the peaked background

A1. List of symbols:

ng: total number of counts in the peak region, obtained from the spectrum

np: number of counts in the peak, obtained from the peak-analysis results

u(np) uncertainty in the number of counts in the peak, obtained from the peak-analysis results

n0: number of counts in the continuum in the peak region

nB: number of counts in the peaked background, obtained from previous measurements

u(nB) uncertainty in the number of counts in the peaked background, obtained from previous measurements

A2. The model

The number of counts in the continuous background is

. 								(A1)

The net number of counts in the peak is

.									(A2)

The uncertainty of the net number of counts is

,							(A3)

because np and nB are obtained from different measurements and are therefore not correlated.

Since in the ROI method n0 and ng are obtained from the spectrum from non-overlapping spectral regions r(ng,n0) = 0; therefore, from

(A4)

it follows

(A5)

A3. Calculation of the uncertainty of the net number of counts as a function of its value:

From Eq. A5 it follows:

.							(A6)

Then, the uncertainty of an arbitrary net number of counts nn’ as a function of its value is:

.		(A7)

Since u2[n0(nn’)] does not depend on the indication, because it was calculated from spectral regions not overlapping with the peak region at the observed value np, it follows:

(A8)

A4. Calculation of the number of counts corresponding to the detection limit and the decision threshold

At nn ‘= 0, the peak-area uncertainty is

(A9)

and the number of counts corresponding to the decision threshold is

.			(A10)

From Eq. (A8) it follows for the uncertainty at the detection limit

.						(A11)

The number of counts corresponding to the detection limit is the solution of Eq. (8), which can be expressed as

(A15)

leading to the quadratic equation  Ann#2 + Bnn# + C = 0, where the coefficients A, B and C assume the values

(A16)

with the solution

.							(A17)